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EM3490/3490B
High Voltage Insulation Resistance Tester
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PRODUCT FUNCTION

[image: EM2600 (1)]· Suitable for measuring the resistance values of various insulating materials and the insulation resistance of transformers, motors and electrical equipment, etc.
· Insulation Resistance Test (IR); Polarization Index Test (PI); Dielectric Absorption Ratio Test (DAR); AC / DC Voltage Test (V).
· 
· 
· The insulation resistance value, test time, and output voltage test data can displays on the same screen. It is very convenient for test operation and historical record query.
· It has unique strong anti-interference technology, with stable testing capability in scenarios with large distributed capacitance (such as long cables) and strong electromagnetic interference environments (such as substations).
· Complete measurement indicators: Insulation Resistance (IR), Polarization Index (PI), Dielectric Absorption Ratio (DAR), AC / DC Voltage (V).
· It is automatic discharge function can automatically and quickly releases the charging charge of the measured object after testing, without an additional discharge circuit.
· The test records can be directly browsed and analyzed on the large-capacity test record memory. It can automatically store 600 groups of real-time test data with test timing.
· The test records can be uploaded to computers for data storage, statistics and analysis by USB communication and upload function.
· The bluetooth communication function can upload the test results to mobile phones.



PRODUCT FEATURES
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	Model
	EM3490
	EM3490B

	Power Supply
	DC11.1V, 2600mAh  rechargeable lithium battery

	Rated Voltage Range
	250V,500V,1000V,2500V,5000V
	500V,1000V,2500V,5000V,10000V

	Output Voltage Accuracy
	+ 4% ± 10V
	+ 4% ± 10V

	Insulation Resistance Measuring Range
	0.1MΩ ~ 5.00TΩ
	0.2MΩ ~ 10.00TΩ

	Short-circuit Current
	5mA MAX
	5mA MAX

	Voltage Measurement
	Range: AC 750V/DC 1000V, Accuracy: ± 1.5%rdg ± 3dgt

	Polarization Index (PI)
	Yes

	Dielectric Absorption Ratio (DAR)
	Yes

	Real-time Voltage Indication
	Real-time monitoring of test terminal voltage; testing is prohibited when the voltage exceeds 36V in non-test state to protect the instrument and operators.

	Measuring Time
	Single test time is 10min; users can stop manually during the test or stop automatically after 10min, and the measurement data is displayed on the screen.

	Storage Function
	Automatically store 600 groups of test data; stored test records can be displayed and played back on the meter.

	USB Communication
	Yes; test records can be transferred to PC via USB cable.

	Bluetooth Communication
	Yes

	Battery Power Indication
	Equipped with battery power indication; when battery voltage is low,it will remind timely charging .

	Auto Power-off Function
	Auto power-off 15 minutes after power-on.

	Protection Level
	IP65 with cover closed; IP40 with cover open

	
	

	
	

	
	

	
	

	
	

	
	

	Power Consumption
	Standby: approx. 40mA (backlight off); Measuring: approx. 100mA (backlight off)

	Working Temperature & Humidity
	- 10℃ ~ 40℃; below 80%RH

	Storage Temperature & Humidity
	- 20℃ ~ 60℃; below 70%RH

	Instrument Dimensions
	280mm × 260mm × 160mm

	Weight
	Approx. 5.58kgs (including accessories)

	Safety Standards
	IEC61010-1 (CAT Ⅲ 300V, CAT IV 150V, Pollution Degree 2)

	Accessories
	Master*1PC
High-voltage test lead (Red)*1PC
High-voltage probe replacement needle*1PC; 
High-voltage probe replacement hook*1PC; 
High-voltage probe replacement alligator clip*1PC; 
Alligator clip test lead (Green)*1PC; 
Alligator clip test lead (Black)*1PC
Charger*1PC
USB communication cable*1PC; 
Tool bag*1PC




Wiring Diagram of Low Power Consumption Module Slave
Power supply & RS485 communication cable
Detector or sensor
Low power consumption module slave
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10th F, U-5, Liandong Intelligent Manufacture Park,           Tel: (86-757)66860936
No. 105, Shilong N. Road, Nanhai, Foshan, China             info@etcr.net
 www.etcrmeter.com

Note! Photographs in this catalogue may slightly differ from the instruments at the time of delivery. 
Subject to technical change without notice.
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