EM3490/3490B

High Voltage Insulation Resistance Tester

Product Function

- Suitable for measuring the resistance values of various insulating materials and the insulation resistance of transformers, motors and electrical equipment, etc.

- Insulation Resistance Test (IR); Polarization Index Test (PI); Dielectric Absorption Ratio Test (DAR); AC / DC Voltage Test (V).

- The insulation resistance value, test time, and output voltage test data can displays on the same screen. It is very convenient for test operation and historical record query.

- It has unique strong anti-interference technology, with stable testing capability in scenarios with large distributed capacitance (such as long cables) and strong electromagnetic interference environments (such as substations).

- Complete measurement indicators: Insulation Resistance (IR), Polarization Index (PI), Dielectric Absorption Ratio (DAR), AC / DC Voltage (V).

- It is automatic discharge function can automatically and quickly releases the charging charge of the measured object after testing, without an additional discharge circuit.

- The test records can be directly browsed and analyzed on the large-capacity test record memory. It can automatically store 600 groups of real-time test data with test timing.

- The test records can be uploaded to computers for data storage, statistics and analysis by USB communication and upload function.

- The bluetooth communication function can upload the test results to mobile phones.

Product Features

Technical Specification

| Model | EM3490 | EM3490B |
| --- | --- | --- |
| Power Supply | DC11.1V, 2600mAh rechargeable lithium battery |  |
| Rated Voltage Range | 250V,500V,1000V,2500V,5000V | 500V,1000V,2500V,5000V,10000V |
| Output Voltage Accuracy | + 4% ± 10V | + 4% ± 10V |
| Insulation Resistance Measuring Range | 0.1MΩ ~ 5.00TΩ | 0.2MΩ ~ 10.00TΩ |
| Short-circuit Current | 5mA MAX | 5mA MAX |
| Voltage Measurement | Range: AC 750V/DC 1000V, Accuracy: ± 1.5%rdg ± 3dgt |  |
| Polarization Index (PI) | Yes |  |
| Dielectric Absorption Ratio (DAR) | Yes |  |
| Real-time Voltage Indication | Real-time monitoring of test terminal voltage; testing is prohibited when the voltage exceeds 36V in non-test state to protect the instrument and operators. |  |
| Measuring Time | Single test time is 10min; users can stop manually during the test or stop automatically after 10min, and the measurement data is displayed on the screen. |  |
| Storage Function | Automatically store 600 groups of test data; stored test records can be displayed and played back on the meter. |  |
| USB Communication | Yes; test records can be transferred to PC via USB cable. |  |
| Bluetooth Communication | Yes |  |
| Battery Power Indication | Equipped with battery power indication; when battery voltage is low,it will remind timely charging . |  |
| Auto Power-off Function | Auto power-off 15 minutes after power-on. |  |
| Protection Level | IP65 with cover closed; IP40 with cover open |  |
| Power Consumption | Standby: approx. 40mA (backlight off); Measuring: approx. 100mA (backlight off) |  |
| Working Temperature & Humidity | - 10℃ ~ 40℃; below 80%RH |  |
| Storage Temperature & Humidity | - 20℃ ~ 60℃; below 70%RH |  |
| Instrument Dimensions | 280mm × 260mm × 160mm |  |
| Weight | Approx. 5.58kgs (including accessories) |  |
| Safety Standards | IEC61010-1 (CAT Ⅲ 300V, CAT IV 150V, Pollution Degree 2) |  |
| Accessories | Master*1PC High-voltage test lead (Red)*1PC High-voltage probe replacement needle*1PC; High-voltage probe replacement hook*1PC; High-voltage probe replacement alligator clip*1PC; Alligator clip test lead (Green)*1PC; Alligator clip test lead (Black)*1PC Charger*1PC USB communication cable*1PC; Tool bag*1PC |  |

Wiring Diagram of Low Power Consumption Module Slave

Power supply & RS485 communication cable

Detector or sensor

Low power consumption module slave

Technical Specification